**Maurice cIEF Cartridge**

|  |  |  |  |
| --- | --- | --- | --- |
| **Part Number** | 090-101 | **Serial Number** |  |
| **Product Name** | Maurice cIEF Cartridge | **Date of Manufacture** |  |
|  |  | **Expiration Date** |  |

This product was tested on a Maurice system with system suitability and passed all physical and functional test criteria.

**Test Methods**

Analysis of Maurice cIEF Cartridge on Maurice or Maurice C. systems. All tests were performed using ProteinSimple’s Maurice cIEF System Suitability. FIGURE 1, representative electropherogram collected in absorbance mode, for pI markers 3.38, 5.85, 6.14, 9.99, and 10.17. FIGURE 2, representative electropherogram collected in fluorescence mode, for the same set of pI markers.

# **ACCEPTANCE RESULTS**

|  |  |  |
| --- | --- | --- |
| PARAMETER | ACCEPTANCE CRITERIA | PASS/FAIL |
| pI Range | pI 3.38, 5.85, 6.14, 9.99, and 10.17 markers observed | Pass |
| Resolution 6.14 (Fluorescence, 10s exposure) | ≥ 2.10 | Pass |
| Resolution 10.17 (Fluorescence, 10s exposure) | ≥ 1.70 | Pass |

Reviewed and Approved:

A picture containing clipart, linedrawing

Description automatically generated

Hooi Tan

Sr. Quality Manager

Date: